## Sheet1
1. Start with Callendar-Van Dusen equation, calculate RTD resistance vs Temperature, over range [-200 degC, 850 degC]
2. Plot the inverse plot, i.e. Temperature vs RTD resistance
3. Use the Excel curve fitting "Trendlines" function, and tell it to create a Trendline using various-ordered polynomials
4. Use Excel trendline fitting coefficients to calculate Temperature vs RTD resistance
5. Calculate a delta-temperature between the curve fitted tempeerature to confirm adequacy of fit
6. Select the desired degree of polynomial fit accuracy, balancing increased computational time (if important to the application)
Fifth order polynomial fit
y = -9.2833E-12x5 + 1.2506E-08x4 - 4.9960E-06x3 + 1.9435E-03x2 + 2.2729E+00x - 2.4297E+02
[10 empty] | R² = 1.0000E+00 | [5 empty] | First order polynomial (i.e. linear) fit
y = 2.8229E+00x - 2.8339E+02
[10 empty] | Fourth order polynomial fit | [5 empty] | R² = 9.9796E-01
Classic Callendar-Van Dusen equations | [9 empty] | y = 2.9401E-09x4 - 1.4136E-06x3 + 1.3543E-03x2 + 2.3132E+00x - 2.4381E+02
R² = 1.0000E+00
RTD(t) = R0*[1 + A*t + B*t^2 + C*(t-100degC)t^3] |  |  | for t<= 0degC
RTD(t) = R0*[1 + A*t + B*t^2] |  |  | for t>= 0degC | (i.e. C = 0) | [5 empty] | Third order polynomial fit
where | [9 empty] | y = 1.0154E-06x3 + 6.8975E-04x2 + 2.3803E+00x - 2.4568E+02
| t = RTD temperature, degC | [8 empty] | R² = 1.0000E+00
| RTD(t) = RTD resistance as a function of RTD temperature (t) |  |  |  |  | from RTD data sheet
| R0 | 100 | R0 = RTD resistance at 100C (usually 100 ohm) |  |  | 0.385Ω/°C |  | 5th order polynomial coeffs |  |  | 4th order polynomial coeffs |  | 3rd order polynomial coeffs |  | 1st order (i.e. linear) coeffs |  | 5th order from AN-709 document, for temps < 0
| A | 0.0039083 | [5 empty] | -9.2833e-12 |  |  | 2.9401e-09 |  | 1.0154e-06 |  | 2.8229 |  | 1.5243e-10
| B | -5.775e-07 |  |  |  | excel linear fit |  | 1.2506e-08 |  |  | -1.4136e-06 |  | 0.00068975 |  | -283.39 |  | -2.8183e-08
| C | -4.183e-12 | for t <= 0 degC |  |  | y = 0.3535x + 100.62 |  | -4.996e-06 |  |  | 0.0013543 |  | 2.3803 |  |  |  | -4.826e-06
|  | 0 | for t >= 0 degC |  |  | R² = 0.998 |  | 0.0019435 |  |  | 2.3132 |  | -245.68 |  |  |  | 0.0025859
[8 empty] | 2.2729 |  |  | -243.81 | [5 empty] | 2.2228
[5 empty] | linear fit coeffs |  |  | -242.97 | [8 empty] | -242.02
0.3535
| Callendar-Van Dusen: Predicting Resistance, given Temp |  |  |  | 100.62 |  |  | Van Dusen-Callendar: Predicting Temp, given Resistance
Temp, degC | Resistance, ohms | A | B | C | fitted ohms | delta ohms |  | fitted temp | delta T |  | fitted temp | delta T | fitted temp | delta T | fitted temp | delta T | fitted temp | delta T
-200 | 18.5201 | 0.0039083 | -5.775e-07 | -4.183e-12 | 29.92 | 11.3999 |  | -200.2394 | -0.23939 |  | -200.5135 | -0.51347 | -201.3536 | -1.3536 | -231.1097 | -31.1097 | -200.0003 | -0.00025913
-195 | 20.6772 | 0.0039083 | -5.775e-07 | -4.183e-12 | 31.6875 | 11.0103 |  | -195.1837 | -0.18372 |  | -195.4124 | -0.41238 | -196.1581 | -1.1581 | -225.0203 | -30.0203 | -195.0003 | -0.00031615
-190 | 22.8255 | 0.0039083 | -5.775e-07 | -4.183e-12 | 33.455 | 10.6295 |  | -190.1335 | -0.13347 |  | -190.3205 | -0.32052 | -190.9771 | -0.97707 | -218.956 | -28.956 | -190.0004 | -0.00035903
-185 | 24.9651 | 0.0039083 | -5.775e-07 | -4.183e-12 | 35.2225 | 10.2574 |  | -185.0884 | -0.088427 |  | -185.2374 | -0.23744 | -185.8098 | -0.80981 | -212.9159 | -27.9159 | -185.0004 | -0.00039126
-180 | 27.0964 | 0.0039083 | -5.775e-07 | -4.183e-12 | 36.99 | 9.8936 |  | -180.0484 | -0.048355 |  | -180.1627 | -0.16272 | -180.6557 | -0.65573 | -206.8995 | -26.8995 | -180.0004 | -0.00041599
-175 | 29.2197 | 0.0039083 | -5.775e-07 | -4.183e-12 | 38.7575 | 9.5378 |  | -175.013 | -0.01303 |  | -175.0959 | -0.095934 | -175.5142 | -0.51423 | -200.9058 | -25.9058 | -175.0004 | -0.00043588
-170 | 31.335 | 0.0039083 | -5.775e-07 | -4.183e-12 | 40.525 | 9.19 |  | -169.9822 | 0.017782 |  | -170.0367 | -0.036663 | -170.3847 | -0.38469 | -194.9343 | -24.9343 | -170.0005 | -0.00045317
-165 | 33.4429 | 0.0039083 | -5.775e-07 | -4.183e-12 | 42.2925 | 8.8496 |  | -164.9557 | 0.044315 |  | -164.9845 | 0.015502 | -165.2666 | -0.26656 | -188.9842 | -23.9842 | -165.0005 | -0.00046961
-160 | 35.5433 | 0.0039083 | -5.775e-07 | -4.183e-12 | 44.06 | 8.5167 |  | -159.9332 | 0.066805 |  | -159.939 | 0.060967 | -160.1592 | -0.15924 | -183.0547 | -23.0547 | -160.0005 | -0.00048649
-155 | 37.6367 | 0.0039083 | -5.775e-07 | -4.183e-12 | 45.8275 | 8.1908 |  | -154.9145 | 0.085489 |  | -154.8999 | 0.10013 | -155.0622 | -0.062199 | -177.1454 | -22.1454 | -155.0005 | -0.00050468
-150 | 39.7232 | 0.0039083 | -5.775e-07 | -4.183e-12 | 47.595 | 7.8718 |  | -149.8994 | 0.1006 |  | -149.8666 | 0.13338 | -149.9749 | 0.02512 | -171.2554 | -21.2554 | -150.0005 | -0.00052469
-145 | 41.803 | 0.0039083 | -5.775e-07 | -4.183e-12 | 49.3625 | 7.5595 |  | -144.8876 | 0.11239 |  | -144.8389 | 0.1611 | -144.8968 | 0.10324 | -165.3842 | -20.3842 | -145.0005 | -0.00054668
-140 | 43.8764 | 0.0039083 | -5.775e-07 | -4.183e-12 | 51.13 | 7.2536 |  | -139.8789 | 0.12107 |  | -139.8163 | 0.18365 | -139.8273 | 0.17269 | -159.5312 | -19.5312 | -140.0006 | -0.00057053
-135 | 45.9436 | 0.0039083 | -5.775e-07 | -4.183e-12 | 52.8975 | 6.9539 |  | -134.8731 | 0.12689 |  | -134.7986 | 0.20142 | -134.766 | 0.23396 | -153.6958 | -18.6958 | -135.0006 | -0.00059591
-130 | 48.0048 | 0.0039083 | -5.775e-07 | -4.183e-12 | 54.665 | 6.6602 |  | -129.8699 | 0.13008 |  | -129.7852 | 0.21476 | -129.7125 | 0.28754 | -147.8774 | -17.8774 | -130.0006 | -0.00062236
-125 | 50.0601 | 0.0039083 | -5.775e-07 | -4.183e-12 | 56.4325 | 6.3724 |  | -124.8691 | 0.13086 |  | -124.776 | 0.224 | -124.6661 | 0.33392 | -142.0754 | -17.0754 | -125.0006 | -0.00064928
-120 | 52.1098 | 0.0039083 | -5.775e-07 | -4.183e-12 | 58.2 | 6.0902 |  | -119.8705 | 0.12946 |  | -119.7705 | 0.2295 | -119.6264 | 0.37355 | -136.2893 | -16.2893 | -120.0007 | -0.00067607
-115 | 54.154 | 0.0039083 | -5.775e-07 | -4.183e-12 | 59.9675 | 5.8135 |  | -114.8739 | 0.1261 |  | -114.7684 | 0.23158 | -114.5931 | 0.40689 | -130.5186 | -15.5186 | -115.0007 | -0.00070213
-110 | 56.193 | 0.0039083 | -5.775e-07 | -4.183e-12 | 61.735 | 5.542 |  | -109.879 | 0.121 |  | -109.7694 | 0.23056 | -109.5656 | 0.43437 | -124.7628 | -14.7628 | -110.0007 | -0.00072691
-105 | 58.2269 | 0.0039083 | -5.775e-07 | -4.183e-12 | 63.5025 | 5.2756 |  | -104.8856 | 0.11435 |  | -104.7732 | 0.22675 | -104.5436 | 0.45642 | -119.0213 | -14.0213 | -105.0007 | -0.00074998
-100 | 60.2558 | 0.0039083 | -5.775e-07 | -4.183e-12 | 65.27 | 5.0142 |  | -99.8936 | 0.10638 |  | -99.7795 | 0.22045 | -99.5266 | 0.47344 | -113.2938 | -13.2938 | -100.0008 | -0.00077103
-95 | 62.28 | 0.0039083 | -5.775e-07 | -4.183e-12 | 67.0375 | 4.7575 |  | -94.9027 | 0.097273 |  | -94.788 | 0.21195 | -94.5142 | 0.48583 | -107.5797 | -12.5797 | -95.0008 | -0.00078989
-90 | 64.2996 | 0.0039083 | -5.775e-07 | -4.183e-12 | 68.805 | 4.5054 |  | -89.9128 | 0.087227 |  | -89.7985 | 0.20153 | -89.506 | 0.49397 | -101.8787 | -11.8787 | -90.0008 | -0.00080659
-85 | 66.3147 | 0.0039083 | -5.775e-07 | -4.183e-12 | 70.5725 | 4.2578 |  | -84.9236 | 0.076426 |  | -84.8105 | 0.18946 | -84.5018 | 0.49823 | -96.1903 | -11.1903 | -85.0008 | -0.0008213
-80 | 68.3254 | 0.0039083 | -5.775e-07 | -4.183e-12 | 72.34 | 4.0146 |  | -79.935 | 0.065048 |  | -79.824 | 0.17598 | -79.501 | 0.49895 | -90.5141 | -10.5141 | -80.0008 | -0.00083437
-75 | 70.332 | 0.0039083 | -5.775e-07 | -4.183e-12 | 74.1075 | 3.7755 |  | -74.9467 | 0.053264 |  | -74.8386 | 0.16135 | -74.5035 | 0.49649 | -84.8497 | -9.8497 | -75.0008 | -0.00084628
-70 | 72.3345 | 0.0039083 | -5.775e-07 | -4.183e-12 | 75.875 | 3.5405 |  | -69.9588 | 0.041238 |  | -69.8542 | 0.14581 | -69.5088 | 0.49116 | -79.1968 | -9.1968 | -70.0009 | -0.00085763
-65 | 74.3331 | 0.0039083 | -5.775e-07 | -4.183e-12 | 77.6425 | 3.3094 |  | -64.9709 | 0.029123 |  | -64.8704 | 0.12956 | -64.5167 | 0.48328 | -73.5551 | -8.5551 | -65.0009 | -0.00086907
-60 | 76.3278 | 0.0039083 | -5.775e-07 | -4.183e-12 | 79.41 | 3.0822 |  | -59.9829 | 0.017064 |  | -59.8872 | 0.11283 | -59.5269 | 0.47314 | -67.9241 | -7.9241 | -60.0009 | -0.00088131
-55 | 78.3189 | 0.0039083 | -5.775e-07 | -4.183e-12 | 81.1775 | 2.8586 |  | -54.9948 | 0.0051999 |  | -54.9042 | 0.095806 | -54.539 | 0.46102 | -62.3037 | -7.3037 | -55.0009 | -0.00089497
-50 | 80.3063 | 0.0039083 | -5.775e-07 | -4.183e-12 | 82.945 | 2.6387 |  | -50.0063 | -0.0063432 |  | -49.9213 | 0.078679 | -49.5528 | 0.44719 | -56.6934 | -6.6934 | -50.0009 | -0.00091056
-45 | 82.2902 | 0.0039083 | -5.775e-07 | -4.183e-12 | 84.7125 | 2.4223 |  | -45.0174 | -0.017447 |  | -44.9384 | 0.061622 | -44.5681 | 0.4319 | -51.0931 | -6.0931 | -45.0009 | -0.0009284
-40 | 84.2707 | 0.0039083 | -5.775e-07 | -4.183e-12 | 86.48 | 2.2093 |  | -40.028 | -0.028003 |  | -39.9552 | 0.044795 | -39.5846 | 0.41539 | -45.5024 | -5.5024 | -40.0009 | -0.00094848
-35 | 86.2478 | 0.0039083 | -5.775e-07 | -4.183e-12 | 88.2475 | 1.9997 |  | -35.0379 | -0.037913 |  | -34.9717 | 0.028346 | -34.6021 | 0.39788 | -39.9211 | -4.9211 | -35.001 | -0.00097039
-30 | 88.2217 | 0.0039083 | -5.775e-07 | -4.183e-12 | 90.015 | 1.7933 |  | -30.0471 | -0.047091 |  | -29.9876 | 0.012409 | -29.6204 | 0.37958 | -34.3491 | -4.3491 | -30.001 | -0.00099322
-25 | 90.1923 | 0.0039083 | -5.775e-07 | -4.183e-12 | 91.7825 | 1.5902 |  | -25.0555 | -0.055459 |  | -25.0029 | -0.0028946 | -24.6393 | 0.36069 | -28.786 | -3.786 | -25.001 | -0.0010154
-20 | 92.1599 | 0.0039083 | -5.775e-07 | -4.183e-12 | 93.55 | 1.3901 |  | -20.063 | -0.06295 |  | -20.0175 | -0.017456 | -19.6586 | 0.34137 | -23.2318 | -3.2318 | -20.001 | -0.0010346
-15 | 94.1244 | 0.0039083 | -5.775e-07 | -4.183e-12 | 95.3175 | 1.1931 |  | -15.0695 | -0.069511 |  | -15.0312 | -0.03118 | -14.6782 | 0.32179 | -17.6862 | -2.6862 | -15.001 | -0.0010475
-10 | 96.0859 | 0.0039083 | -5.775e-07 | -4.183e-12 | 97.085 | 0.99912 |  | -10.0751 | -0.075097 |  | -10.044 | -0.043986 | -9.6979 | 0.30211 | -12.1492 | -2.1492 | -10.001 | -0.0010499
-5 | 98.0444 | 0.0039083 | -5.775e-07 | -4.183e-12 | 98.8525 | 0.8081 |  | -5.0797 | -0.079678 |  | -5.0558 | -0.055807 | -4.7176 | 0.28244 | -6.6205 | -1.6205 | -5.001 | -0.0010364
0 | 100 | 0.0039083 | -5.775e-07 | 0 | 100.62 | 0.62 |  | -0.083233 | -0.083233 |  | -0.06659 | -0.06659 | 0.2629 | 0.2629 | -1.1 | -1.1 | -0.001 | -0.001
5 | 101.9527 | 0.0039083 | -5.775e-07 | 0 | 102.3875 | 0.43479 |  | 4.9142 | -0.085767 |  | 4.9237 | -0.076308 | 5.2436 | 0.24359 | 4.4123 | -0.58771 | 4.9991 | -0.0009453
10 | 103.9025 | 0.0039083 | -5.775e-07 | 0 | 104.155 | 0.25248 |  | 9.9127 | -0.087344 |  | 9.915 | -0.084994 | 10.2245 | 0.22452 | 9.9164 | -0.083562 | 9.9991 | -0.0009208
15 | 105.8495 | 0.0039083 | -5.775e-07 | 0 | 105.9225 | 0.073044 |  | 14.912 | -0.088036 |  | 14.9073 | -0.092696 | 15.2057 | 0.20571 | 15.4124 | 0.41243 | 14.999 | -0.00097158
20 | 107.7935 | 0.0039083 | -5.775e-07 | 0 | 107.69 | -0.1035 |  | 19.9121 | -0.087916 |  | 19.9005 | -0.099457 | 20.1872 | 0.18716 | 20.9003 | 0.90027 | 19.9989 | -0.001125
25 | 109.7347 | 0.0039083 | -5.775e-07 | 0 | 109.4575 | -0.27716 |  | 24.9129 | -0.08705 |  | 24.8947 | -0.10532 | 25.1689 | 0.1689 | 26.38 | 1.38 | 24.9986 | -0.0013905
30 | 111.6729 | 0.0039083 | -5.775e-07 | 0 | 111.225 | -0.44792 |  | 29.9145 | -0.085505 |  | 29.8897 | -0.11034 | 30.1509 | 0.15093 | 31.8515 | 1.8515 | 29.9982 | -0.0017593
35 | 113.6083 | 0.0039083 | -5.775e-07 | 0 | 112.9925 | -0.61581 |  | 34.9167 | -0.083342 |  | 34.8855 | -0.11454 | 35.1333 | 0.13326 | 37.3149 | 2.3149 | 34.9978 | -0.0022041
40 | 115.5408 | 0.0039083 | -5.775e-07 | 0 | 114.76 | -0.7808 |  | 39.9194 | -0.080621 |  | 39.882 | -0.11797 | 40.1159 | 0.11589 | 42.7701 | 2.7701 | 39.9973 | -0.002679
45 | 117.4704 | 0.0039083 | -5.775e-07 | 0 | 116.5275 | -0.94291 |  | 44.9226 | -0.077398 |  | 44.8793 | -0.12067 | 45.0988 | 0.098845 | 48.2172 | 3.2172 | 44.9969 | -0.0031194
50 | 119.3971 | 0.0039083 | -5.775e-07 | 0 | 118.295 | -1.1021 |  | 49.9263 | -0.073728 |  | 49.8773 | -0.12268 | 50.0821 | 0.082126 | 53.6561 | 3.6561 | 49.9966 | -0.0034413
55 | 121.321 | 0.0039083 | -5.775e-07 | 0 | 120.0625 | -1.2585 |  | 54.9303 | -0.069663 |  | 54.876 | -0.12404 | 55.0657 | 0.065742 | 59.0869 | 4.0869 | 54.9965 | -0.0035417
60 | 123.2419 | 0.0039083 | -5.775e-07 | 0 | 121.83 | -1.4119 |  | 59.9347 | -0.065251 |  | 59.8752 | -0.12478 | 60.0497 | 0.049702 | 64.5096 | 4.5096 | 59.9967 | -0.0032982
65 | 125.16 | 0.0039083 | -5.775e-07 | 0 | 123.5975 | -1.5625 |  | 64.9395 | -0.06054 |  | 64.8751 | -0.12493 | 65.034 | 0.034014 | 69.924 | 4.924 | 64.9974 | -0.0025685
70 | 127.0751 | 0.0039083 | -5.775e-07 | 0 | 125.365 | -1.7101 |  | 69.9444 | -0.055573 |  | 69.8755 | -0.12454 | 70.0187 | 0.018685 | 75.3304 | 5.3304 | 69.9988 | -0.0011909
75 | 128.9874 | 0.0039083 | -5.775e-07 | 0 | 127.1325 | -1.8549 |  | 74.9496 | -0.050393 |  | 74.8764 | -0.12363 | 75.0037 | 0.0037215 | 80.7285 | 5.7285 | 75.001 | 0.0010166
80 | 130.8968 | 0.0039083 | -5.775e-07 | 0 | 128.9 | -1.9968 |  | 79.955 | -0.04504 |  | 79.8778 | -0.12225 | 79.9891 | -0.010869 | 86.1186 | 6.1186 | 80.0043 | 0.0042558
85 | 132.8033 | 0.0039083 | -5.775e-07 | 0 | 130.6675 | -2.1358 |  | 84.9604 | -0.039552 |  | 84.8796 | -0.12041 | 84.9749 | -0.025079 | 91.5005 | 6.5005 | 85.0087 | 0.0087488
90 | 134.7069 | 0.0039083 | -5.775e-07 | 0 | 132.435 | -2.2719 |  | 89.966 | -0.033963 |  | 89.8818 | -0.11815 | 89.9611 | -0.038904 | 96.8742 | 6.8742 | 90.0147 | 0.014738
95 | 136.6077 | 0.0039083 | -5.775e-07 | 0 | 134.2025 | -2.4052 |  | 94.9717 | -0.028307 |  | 94.8845 | -0.11551 | 94.9477 | -0.052337 | 102.2398 | 7.2398 | 95.0225 | 0.022487
100 | 138.5055 | 0.0039083 | -5.775e-07 | 0 | 135.97 | -2.5355 |  | 99.9774 | -0.022617 |  | 99.8875 | -0.1125 | 99.9346 | -0.065374 | 107.5972 | 7.5972 | 100.0323 | 0.032278
105 | 140.4005 | 0.0039083 | -5.775e-07 | 0 | 137.7375 | -2.663 |  | 104.9831 | -0.016921 |  | 104.8908 | -0.10916 | 104.922 | -0.078008 | 112.9464 | 7.9464 | 105.0444 | 0.044417
110 | 142.2925 | 0.0039083 | -5.775e-07 | 0 | 139.505 | -2.7875 |  | 109.9888 | -0.011248 |  | 109.8945 | -0.10552 | 109.9098 | -0.090235 | 118.2876 | 8.2876 | 110.0592 | 0.059227
115 | 144.1817 | 0.0039083 | -5.775e-07 | 0 | 141.2725 | -2.9092 |  | 114.9944 | -0.0056227 |  | 114.8984 | -0.1016 | 114.898 | -0.10205 | 123.6205 | 8.6205 | 115.0771 | 0.077056
120 | 146.068 | 0.0039083 | -5.775e-07 | 0 | 143.04 | -3.028 |  | 119.9999 | -6.9746e-05 |  | 119.9026 | -0.097422 | 119.8866 | -0.11345 | 128.9454 | 8.9454 | 120.0983 | 0.098269
125 | 147.9514 | 0.0039083 | -5.775e-07 | 0 | 144.8075 | -3.1439 |  | 125.0054 | 0.0053886 |  | 124.907 | -0.093017 | 124.8756 | -0.12443 | 134.262 | 9.262 | 125.1233 | 0.12326
130 | 149.8319 | 0.0039083 | -5.775e-07 | 0 | 146.575 | -3.2569 |  | 130.0107 | 0.010731 |  | 129.9116 | -0.088405 | 129.865 | -0.13499 | 139.5705 | 9.5705 | 130.1524 | 0.15242
135 | 151.7096 | 0.0039083 | -5.775e-07 | 0 | 148.3425 | -3.3671 |  | 135.0159 | 0.01594 |  | 134.9164 | -0.08361 | 134.8549 | -0.14511 | 144.8709 | 9.8709 | 135.1862 | 0.1862
140 | 153.5843 | 0.0039083 | -5.775e-07 | 0 | 150.11 | -3.4743 |  | 140.021 | 0.020996 |  | 139.9213 | -0.078655 | 139.8452 | -0.15481 | 150.1631 | 10.1631 | 140.225 | 0.22505
145 | 155.4562 | 0.0039083 | -5.775e-07 | 0 | 151.8775 | -3.5787 |  | 145.0259 | 0.025885 |  | 144.9264 | -0.07356 | 144.8359 | -0.16408 | 155.4472 | 10.4472 | 145.2694 | 0.26943
150 | 157.3251 | 0.0039083 | -5.775e-07 | 0 | 153.645 | -3.6801 |  | 150.0306 | 0.03059 |  | 149.9317 | -0.068346 | 149.8271 | -0.17292 | 160.7231 | 10.7231 | 150.3198 | 0.31984
155 | 159.1912 | 0.0039083 | -5.775e-07 | 0 | 155.4125 | -3.7787 |  | 155.0351 | 0.0351 |  | 154.937 | -0.063032 | 154.8187 | -0.18131 | 165.9909 | 10.9909 | 155.3768 | 0.37679
160 | 161.0544 | 0.0039083 | -5.775e-07 | 0 | 157.18 | -3.8744 |  | 160.0394 | 0.039403 |  | 159.9424 | -0.057637 | 159.8107 | -0.18927 | 171.2505 | 11.2505 | 160.4408 | 0.44083
165 | 162.9147 | 0.0039083 | -5.775e-07 | 0 | 158.9475 | -3.9672 |  | 165.0435 | 0.043487 |  | 164.9478 | -0.05218 | 164.8032 | -0.19679 | 176.5019 | 11.5019 | 165.5125 | 0.51251
170 | 164.7721 | 0.0039083 | -5.775e-07 | 0 | 160.715 | -4.0571 |  | 170.0473 | 0.047344 |  | 169.9533 | -0.046679 | 169.7961 | -0.20387 | 181.7452 | 11.7452 | 170.5924 | 0.59241
175 | 166.6267 | 0.0039083 | -5.775e-07 | 0 | 162.4825 | -4.1442 |  | 175.051 | 0.050966 |  | 174.9589 | -0.041148 | 174.7895 | -0.2105 | 186.9804 | 11.9804 | 175.6811 | 0.68114
180 | 168.4783 | 0.0039083 | -5.775e-07 | 0 | 164.25 | -4.2283 |  | 180.0543 | 0.054346 |  | 179.9644 | -0.035606 | 179.7833 | -0.2167 | 192.2074 | 12.2074 | 180.7793 | 0.7793
185 | 170.3271 | 0.0039083 | -5.775e-07 | 0 | 166.0175 | -4.3096 |  | 185.0575 | 0.057478 |  | 184.9699 | -0.030067 | 184.7775 | -0.22245 | 197.4262 | 12.4262 | 185.8876 | 0.88756
190 | 172.1729 | 0.0039083 | -5.775e-07 | 0 | 167.785 | -4.3879 |  | 190.0604 | 0.060358 |  | 189.9755 | -0.024545 | 189.7722 | -0.22776 | 202.6369 | 12.6369 | 191.0066 | 1.0066
195 | 174.0159 | 0.0039083 | -5.775e-07 | 0 | 169.5525 | -4.4634 |  | 195.063 | 0.062982 |  | 194.9809 | -0.019056 | 194.7674 | -0.23263 | 207.8395 | 12.8395 | 196.137 | 1.137
200 | 175.856 | 0.0039083 | -5.775e-07 | 0 | 171.32 | -4.536 |  | 200.0653 | 0.065347 |  | 199.9864 | -0.013611 | 199.7629 | -0.23706 | 213.0339 | 13.0339 | 201.2796 | 1.2796
205 | 177.6932 | 0.0039083 | -5.775e-07 | 0 | 173.0875 | -4.6057 |  | 205.0675 | 0.067453 |  | 204.9918 | -0.0082237 | 204.759 | -0.24105 | 218.2202 | 13.2202 | 206.4351 | 1.4351
210 | 179.5275 | 0.0039083 | -5.775e-07 | 0 | 174.855 | -4.6725 |  | 210.0693 | 0.069298 |  | 209.9971 | -0.0029063 | 209.7554 | -0.2446 | 223.3983 | 13.3983 | 211.6042 | 1.6042
215 | 181.359 | 0.0039083 | -5.775e-07 | 0 | 176.6225 | -4.7365 |  | 215.0709 | 0.070883 |  | 215.0023 | 0.0023299 | 214.7523 | -0.24771 | 228.5682 | 13.5682 | 216.7877 | 1.7877
220 | 183.1875 | 0.0039083 | -5.775e-07 | 0 | 178.39 | -4.7975 |  | 220.0722 | 0.072209 |  | 220.0075 | 0.0074741 | 219.7496 | -0.25038 | 233.73 | 13.73 | 221.9864 | 1.9864
225 | 185.0132 | 0.0039083 | -5.775e-07 | 0 | 180.1575 | -4.8557 |  | 225.0733 | 0.073277 |  | 225.0125 | 0.012516 | 224.7474 | -0.25263 | 238.8836 | 13.8836 | 227.2011 | 2.2011
230 | 186.8359 | 0.0039083 | -5.775e-07 | 0 | 181.925 | -4.9109 |  | 230.0741 | 0.07409 |  | 230.0174 | 0.017447 | 229.7456 | -0.25444 | 244.0291 | 14.0291 | 232.4326 | 2.4326
235 | 188.6558 | 0.0039083 | -5.775e-07 | 0 | 183.6925 | -4.9633 |  | 235.0747 | 0.074652 |  | 235.0223 | 0.022256 | 234.7442 | -0.25583 | 249.1665 | 14.1665 | 237.6818 | 2.6818
240 | 190.4728 | 0.0039083 | -5.775e-07 | 0 | 185.46 | -5.0128 |  | 240.075 | 0.074967 |  | 240.0269 | 0.026936 | 239.7432 | -0.25679 | 254.2957 | 14.2957 | 242.9496 | 2.9496
245 | 192.2869 | 0.0039083 | -5.775e-07 | 0 | 187.2275 | -5.0594 |  | 245.075 | 0.075039 |  | 245.0315 | 0.031479 | 244.7427 | -0.25733 | 259.4167 | 14.4167 | 248.2369 | 3.2369
250 | 194.0981 | 0.0039083 | -5.775e-07 | 0 | 188.995 | -5.1031 |  | 250.0749 | 0.074874 |  | 250.0359 | 0.035878 | 249.7425 | -0.25745 | 264.5296 | 14.5296 | 253.5446 | 3.5446
255 | 195.9065 | 0.0039083 | -5.775e-07 | 0 | 190.7625 | -5.144 |  | 255.0745 | 0.074477 |  | 255.0401 | 0.040124 | 254.7428 | -0.25716 | 269.6343 | 14.6343 | 258.8736 | 3.8736
260 | 197.7119 | 0.0039083 | -5.775e-07 | 0 | 192.53 | -5.1819 |  | 260.0739 | 0.073855 |  | 260.0442 | 0.044214 | 259.7435 | -0.25646 | 274.7309 | 14.7309 | 264.2249 | 4.2249
265 | 199.5145 | 0.0039083 | -5.775e-07 | 0 | 194.2975 | -5.217 |  | 265.073 | 0.073016 |  | 265.0481 | 0.048139 | 264.7446 | -0.25536 | 279.8194 | 14.8194 | 269.5995 | 4.5995
270 | 201.3141 | 0.0039083 | -5.775e-07 | 0 | 196.065 | -5.2491 |  | 270.072 | 0.071966 |  | 270.0519 | 0.051896 | 269.7461 | -0.25385 | 284.8996 | 14.8996 | 274.9983 | 4.9983
275 | 203.1109 | 0.0039083 | -5.775e-07 | 0 | 197.8325 | -5.2784 |  | 275.0707 | 0.070713 |  | 275.0555 | 0.055479 | 274.748 | -0.25195 | 289.9718 | 14.9718 | 280.4225 | 5.4225
280 | 204.9048 | 0.0039083 | -5.775e-07 | 0 | 199.6 | -5.3048 |  | 280.0693 | 0.069265 |  | 280.0589 | 0.058884 | 279.7503 | -0.24966 | 295.0358 | 15.0358 | 285.873 | 5.873
285 | 206.6958 | 0.0039083 | -5.775e-07 | 0 | 201.3675 | -5.3283 |  | 285.0676 | 0.067632 |  | 285.0621 | 0.062107 | 284.753 | -0.24699 | 300.0916 | 15.0916 | 291.3508 | 6.3508
290 | 208.4839 | 0.0039083 | -5.775e-07 | 0 | 203.135 | -5.3489 |  | 290.0658 | 0.065822 |  | 290.0651 | 0.065144 | 289.7561 | -0.24393 | 305.1393 | 15.1393 | 296.8572 | 6.8572
295 | 210.2692 | 0.0039083 | -5.775e-07 | 0 | 204.9025 | -5.3667 |  | 295.0638 | 0.063845 |  | 295.068 | 0.067993 | 294.7595 | -0.24051 | 310.1788 | 15.1788 | 302.3931 | 7.3931
300 | 212.0515 | 0.0039083 | -5.775e-07 | 0 | 206.67 | -5.3815 |  | 300.0617 | 0.06171 |  | 300.0707 | 0.070652 | 299.7633 | -0.23671 | 315.2102 | 15.2102 | 307.9598 | 7.9598
305 | 213.831 | 0.0039083 | -5.775e-07 | 0 | 208.4375 | -5.3935 |  | 305.0594 | 0.059427 |  | 305.0731 | 0.073117 | 304.7674 | -0.23256 | 320.2334 | 15.2334 | 313.5582 | 8.5582
310 | 215.6075 | 0.0039083 | -5.775e-07 | 0 | 210.205 | -5.4025 |  | 310.057 | 0.057007 |  | 310.0754 | 0.075387 | 309.7719 | -0.22806 | 325.2485 | 15.2485 | 319.1897 | 9.1897
315 | 217.3812 | 0.0039083 | -5.775e-07 | 0 | 211.9725 | -5.4087 |  | 315.0545 | 0.054459 |  | 315.0775 | 0.077462 | 314.7768 | -0.22321 | 330.2554 | 15.2554 | 324.8554 | 9.8554
320 | 219.152 | 0.0039083 | -5.775e-07 | 0 | 213.74 | -5.412 |  | 320.0518 | 0.051794 |  | 320.0793 | 0.079339 | 319.782 | -0.21802 | 335.2542 | 15.2542 | 330.5564 | 10.5564
325 | 220.9199 | 0.0039083 | -5.775e-07 | 0 | 215.5075 | -5.4124 |  | 325.049 | 0.049023 |  | 325.081 | 0.08102 | 324.7875 | -0.21251 | 340.2448 | 15.2448 | 336.2941 | 11.2941
330 | 222.6849 | 0.0039083 | -5.775e-07 | 0 | 217.275 | -5.4099 |  | 330.0462 | 0.046156 |  | 330.0825 | 0.082502 | 329.7933 | -0.20667 | 345.2273 | 15.2273 | 342.0695 | 12.0695
335 | 224.4471 | 0.0039083 | -5.775e-07 | 0 | 219.0425 | -5.4046 |  | 335.0432 | 0.043205 |  | 335.0838 | 0.083788 | 334.7995 | -0.20052 | 350.2016 | 15.2016 | 347.8841 | 12.8841
340 | 226.2063 | 0.0039083 | -5.775e-07 | 0 | 220.81 | -5.3963 |  | 340.0402 | 0.04018 |  | 340.0849 | 0.084876 | 339.8059 | -0.19406 | 355.1678 | 15.1678 | 353.739 | 13.739
345 | 227.9627 | 0.0039083 | -5.775e-07 | 0 | 222.5775 | -5.3852 |  | 345.0371 | 0.037093 |  | 345.0858 | 0.085769 | 344.8127 | -0.1873 | 360.1258 | 15.1258 | 359.6356 | 14.6356
350 | 229.7161 | 0.0039083 | -5.775e-07 | 0 | 224.345 | -5.3711 |  | 350.034 | 0.033954 |  | 350.0865 | 0.086467 | 349.8197 | -0.18026 | 365.0756 | 15.0756 | 365.5751 | 15.5751
355 | 231.4667 | 0.0039083 | -5.775e-07 | 0 | 226.1125 | -5.3542 |  | 355.0308 | 0.030774 |  | 355.087 | 0.086972 | 354.8271 | -0.17294 | 370.0174 | 15.0174 | 371.559 | 16.559
360 | 233.2144 | 0.0039083 | -5.775e-07 | 0 | 227.88 | -5.3344 |  | 360.0276 | 0.027565 |  | 360.0873 | 0.087286 | 359.8346 | -0.16536 | 374.9509 | 14.9509 | 377.5884 | 17.5884
365 | 234.9592 | 0.0039083 | -5.775e-07 | 0 | 229.6475 | -5.3117 |  | 365.0243 | 0.024336 |  | 365.0874 | 0.087412 | 364.8425 | -0.15751 | 379.8763 | 14.8763 | 383.6649 | 18.6649
370 | 236.7011 | 0.0039083 | -5.775e-07 | 0 | 231.415 | -5.2861 |  | 370.0211 | 0.0211 |  | 370.0874 | 0.087351 | 369.8506 | -0.14941 | 384.7936 | 14.7936 | 389.7897 | 19.7897
375 | 238.4402 | 0.0039083 | -5.775e-07 | 0 | 233.1825 | -5.2577 |  | 375.0179 | 0.017867 |  | 375.0871 | 0.087107 | 374.8589 | -0.14108 | 389.7027 | 14.7027 | 395.9643 | 20.9643
380 | 240.1763 | 0.0039083 | -5.775e-07 | 0 | 234.95 | -5.2263 |  | 380.0146 | 0.014648 |  | 380.0867 | 0.086683 | 379.8675 | -0.13252 | 394.6037 | 14.6037 | 402.1901 | 22.1901
385 | 241.9096 | 0.0039083 | -5.775e-07 | 0 | 236.7175 | -5.1921 |  | 385.0115 | 0.011452 |  | 385.0861 | 0.086081 | 384.8763 | -0.12374 | 399.4965 | 14.4965 | 408.4686 | 23.4686
390 | 243.6399 | 0.0039083 | -5.775e-07 | 0 | 238.485 | -5.1549 |  | 390.0083 | 0.0082911 |  | 390.0853 | 0.085307 | 389.8852 | -0.11475 | 404.3811 | 14.3811 | 414.8011 | 24.8011
395 | 245.3674 | 0.0039083 | -5.775e-07 | 0 | 240.2525 | -5.1149 |  | 395.0052 | 0.0051747 |  | 395.0844 | 0.084363 | 394.8944 | -0.10557 | 409.2577 | 14.2577 | 421.1891 | 26.1891
400 | 247.092 | 0.0039083 | -5.775e-07 | 0 | 242.02 | -5.072 |  | 400.0021 | 0.0021129 |  | 400.0833 | 0.083253 | 399.9038 | -0.096208 | 414.126 | 14.126 | 427.6341 | 27.6341
405 | 248.8137 | 0.0039083 | -5.775e-07 | 0 | 243.7875 | -5.0262 |  | 404.9991 | -0.0008845 |  | 405.082 | 0.081983 | 404.9133 | -0.086673 | 418.9862 | 13.9862 | 434.1377 | 29.1377
410 | 250.5325 | 0.0039083 | -5.775e-07 | 0 | 245.555 | -4.9775 |  | 409.9962 | -0.0038079 |  | 410.0806 | 0.080557 | 409.923 | -0.076978 | 423.8383 | 13.8383 | 440.7013 | 30.7013
415 | 252.2485 | 0.0039083 | -5.775e-07 | 0 | 247.3225 | -4.926 |  | 414.9934 | -0.0066481 |  | 415.079 | 0.07898 | 414.9329 | -0.067137 | 428.6822 | 13.6822 | 447.3265 | 32.3265
420 | 253.9615 | 0.0039083 | -5.775e-07 | 0 | 249.09 | -4.8715 |  | 419.9906 | -0.009396 |  | 420.0773 | 0.077256 | 419.9428 | -0.05716 | 433.5179 | 13.5179 | 454.0147 | 34.0147
425 | 255.6717 | 0.0039083 | -5.775e-07 | 0 | 250.8575 | -4.8142 |  | 424.988 | -0.012043 |  | 425.0754 | 0.075391 | 424.9529 | -0.047062 | 438.3455 | 13.3455 | 460.7676 | 35.7676
430 | 257.3789 | 0.0039083 | -5.775e-07 | 0 | 252.625 | -4.7539 |  | 429.9854 | -0.01458 |  | 430.0734 | 0.07339 | 429.9631 | -0.036854 | 443.165 | 13.165 | 467.5868 | 37.5868
435 | 259.0833 | 0.0039083 | -5.775e-07 | 0 | 254.3925 | -4.6908 |  | 434.983 | -0.017 |  | 435.0713 | 0.07126 | 434.9734 | -0.02655 | 447.9763 | 12.9763 | 474.4737 | 39.4737
440 | 260.7848 | 0.0039083 | -5.775e-07 | 0 | 256.16 | -4.6248 |  | 439.9807 | -0.019295 |  | 440.069 | 0.069006 | 439.9838 | -0.016164 | 452.7794 | 12.7794 | 481.4301 | 41.4301
445 | 262.4834 | 0.0039083 | -5.775e-07 | 0 | 257.9275 | -4.5559 |  | 444.9785 | -0.021457 |  | 445.0666 | 0.066633 | 444.9943 | -0.0057096 | 457.5744 | 12.5744 | 488.4575 | 43.4575
450 | 264.1791 | 0.0039083 | -5.775e-07 | 0 | 259.695 | -4.4841 |  | 449.9765 | -0.023478 |  | 450.0641 | 0.064149 | 450.0048 | 0.0048001 | 462.3613 | 12.3613 | 495.5576 | 45.5576
455 | 265.872 | 0.0039083 | -5.775e-07 | 0 | 261.4625 | -4.4095 |  | 454.9746 | -0.025354 |  | 455.0616 | 0.061559 | 455.0154 | 0.015351 | 467.1399 | 12.1399 | 502.7321 | 47.7321
460 | 267.5619 | 0.0039083 | -5.775e-07 | 0 | 263.23 | -4.3319 |  | 459.9729 | -0.027076 |  | 460.0589 | 0.05887 | 460.0259 | 0.025927 | 471.9105 | 11.9105 | 509.9825 | 49.9825
465 | 269.249 | 0.0039083 | -5.775e-07 | 0 | 264.9975 | -4.2515 |  | 464.9714 | -0.02864 |  | 465.0561 | 0.056088 | 465.0365 | 0.036516 | 476.6729 | 11.6729 | 517.3105 | 52.3105
470 | 270.9331 | 0.0039083 | -5.775e-07 | 0 | 266.765 | -4.1681 |  | 469.97 | -0.03004 |  | 470.0532 | 0.053221 | 470.0471 | 0.047102 | 481.4271 | 11.4271 | 524.7179 | 54.7179
475 | 272.6144 | 0.0039083 | -5.775e-07 | 0 | 268.5325 | -4.0819 |  | 474.9687 | -0.031271 |  | 475.0503 | 0.050275 | 475.0577 | 0.05767 | 486.1732 | 11.1732 | 532.2063 | 57.2063
480 | 274.2928 | 0.0039083 | -5.775e-07 | 0 | 270.3 | -3.9928 |  | 479.9677 | -0.032328 |  | 480.0473 | 0.047257 | 480.0682 | 0.068204 | 490.9111 | 10.9111 | 539.7775 | 59.7775
485 | 275.9683 | 0.0039083 | -5.775e-07 | 0 | 272.0675 | -3.9008 |  | 484.9668 | -0.033207 |  | 485.0442 | 0.044175 | 485.0787 | 0.07869 | 495.6409 | 10.6409 | 547.4331 | 62.4331
490 | 277.6409 | 0.0039083 | -5.775e-07 | 0 | 273.835 | -3.8059 |  | 489.9661 | -0.033904 |  | 490.041 | 0.041034 | 490.0891 | 0.089111 | 500.3626 | 10.3626 | 555.175 | 65.175
495 | 279.3107 | 0.0039083 | -5.775e-07 | 0 | 275.6025 | -3.7082 |  | 494.9656 | -0.034417 |  | 495.0378 | 0.037844 | 495.0995 | 0.099452 | 505.0761 | 10.0761 | 563.0048 | 68.0048
500 | 280.9775 | 0.0039083 | -5.775e-07 | 0 | 277.37 | -3.6075 |  | 499.9653 | -0.034743 |  | 500.0346 | 0.034611 | 500.1097 | 0.1097 | 509.7814 | 9.7814 | 570.9243 | 70.9243
505 | 282.6415 | 0.0039083 | -5.775e-07 | 0 | 279.1375 | -3.504 |  | 504.9651 | -0.03488 |  | 505.0313 | 0.031343 | 505.1198 | 0.11983 | 514.4786 | 9.4786 | 578.9353 | 73.9353
510 | 284.3025 | 0.0039083 | -5.775e-07 | 0 | 280.905 | -3.3975 |  | 509.9652 | -0.034825 |  | 510.028 | 0.028046 | 510.1298 | 0.12983 | 519.1676 | 9.1676 | 587.0395 | 77.0395
515 | 285.9607 | 0.0039083 | -5.775e-07 | 0 | 282.6725 | -3.2882 |  | 514.9654 | -0.034578 |  | 515.0247 | 0.02473 | 515.1397 | 0.13969 | 523.8485 | 8.8485 | 595.2388 | 80.2388
520 | 287.616 | 0.0039083 | -5.775e-07 | 0 | 284.44 | -3.176 |  | 519.9659 | -0.034139 |  | 520.0214 | 0.021401 | 520.1494 | 0.14938 | 528.5212 | 8.5212 | 603.535 | 83.535
525 | 289.2684 | 0.0039083 | -5.775e-07 | 0 | 286.2075 | -3.0609 |  | 524.9665 | -0.033506 |  | 525.0181 | 0.018068 | 525.1589 | 0.15889 | 533.1858 | 8.1858 | 611.9298 | 86.9298
530 | 290.9179 | 0.0039083 | -5.775e-07 | 0 | 287.975 | -2.9429 |  | 529.9673 | -0.032682 |  | 530.0147 | 0.014738 | 530.1682 | 0.16821 | 537.8422 | 7.8422 | 620.4251 | 90.4251
535 | 292.5646 | 0.0039083 | -5.775e-07 | 0 | 289.7425 | -2.8221 |  | 534.9683 | -0.031667 |  | 535.0114 | 0.011418 | 535.1773 | 0.17731 | 542.4905 | 7.4905 | 629.0228 | 94.0228
540 | 294.2083 | 0.0039083 | -5.775e-07 | 0 | 291.51 | -2.6983 |  | 539.9695 | -0.030462 |  | 540.0081 | 0.0081176 | 540.1862 | 0.18617 | 547.1306 | 7.1306 | 637.7246 | 97.7246
545 | 295.8492 | 0.0039083 | -5.775e-07 | 0 | 293.2775 | -2.5717 |  | 544.9709 | -0.02907 |  | 545.0048 | 0.0048437 | 545.1948 | 0.19479 | 551.7626 | 6.7626 | 646.5325 | 101.5325
550 | 297.4871 | 0.0039083 | -5.775e-07 | 0 | 295.045 | -2.4421 |  | 549.9725 | -0.027494 |  | 550.0016 | 0.0016044 | 550.2031 | 0.20314 | 556.3864 | 6.3864 | 655.4482 | 105.4482
555 | 299.1222 | 0.0039083 | -5.775e-07 | 0 | 296.8125 | -2.3097 |  | 554.9743 | -0.025737 |  | 554.9984 | -0.0015923 | 555.2112 | 0.21119 | 561.0021 | 6.0021 | 664.4737 | 109.4737
560 | 300.7544 | 0.0039083 | -5.775e-07 | 0 | 298.58 | -2.1744 |  | 559.9762 | -0.023804 |  | 559.9953 | -0.0047385 | 560.2189 | 0.21894 | 565.6096 | 5.6096 | 673.6108 | 113.6108
565 | 302.3837 | 0.0039083 | -5.775e-07 | 0 | 300.3475 | -2.0362 |  | 564.9783 | -0.021698 |  | 564.9922 | -0.0078263 | 565.2264 | 0.22637 | 570.209 | 5.209 | 682.8614 | 117.8614
570 | 304.0101 | 0.0039083 | -5.775e-07 | 0 | 302.115 | -1.8951 |  | 569.9806 | -0.019426 |  | 569.9892 | -0.010848 | 570.2335 | 0.23345 | 574.8002 | 4.8002 | 692.2274 | 122.2274
575 | 305.6337 | 0.0039083 | -5.775e-07 | 0 | 303.8825 | -1.7512 |  | 574.983 | -0.016992 |  | 574.9862 | -0.013795 | 575.2402 | 0.24017 | 579.3832 | 4.3832 | 701.7108 | 126.7108
580 | 307.2543 | 0.0039083 | -5.775e-07 | 0 | 305.65 | -1.6043 |  | 579.9856 | -0.014405 |  | 579.9833 | -0.01666 | 580.2465 | 0.2465 | 583.9582 | 3.9582 | 711.3134 | 131.3134
585 | 308.8721 | 0.0039083 | -5.775e-07 | 0 | 307.4175 | -1.4546 |  | 584.9883 | -0.01167 |  | 584.9806 | -0.019435 | 585.2524 | 0.25244 | 588.5249 | 3.5249 | 721.0371 | 136.0371
590 | 310.4869 | 0.0039083 | -5.775e-07 | 0 | 309.185 | -1.3019 |  | 589.9912 | -0.0087961 |  | 589.9779 | -0.022111 | 590.2579 | 0.25794 | 593.0835 | 3.0835 | 730.8838 | 140.8838
595 | 312.0989 | 0.0039083 | -5.775e-07 | 0 | 310.9525 | -1.1464 |  | 594.9942 | -0.0057913 |  | 594.9753 | -0.024683 | 595.263 | 0.26301 | 597.634 | 2.634 | 740.8556 | 145.8556
600 | 313.708 | 0.0039083 | -5.775e-07 | 0 | 312.72 | -0.988 |  | 599.9973 | -0.0026648 |  | 599.9729 | -0.02714 | 600.2676 | 0.26761 | 602.1763 | 2.1763 | 750.9543 | 150.9543
605 | 315.3142 | 0.0039083 | -5.775e-07 | 0 | 314.4875 | -0.82671 |  | 605.0006 | 0.0005736 |  | 604.9705 | -0.029476 | 605.2717 | 0.27173 | 606.7105 | 1.7105 | 761.1819 | 156.1819
610 | 316.9175 | 0.0039083 | -5.775e-07 | 0 | 316.255 | -0.66253 |  | 610.0039 | 0.0039135 |  | 609.9683 | -0.031683 | 610.2754 | 0.27535 | 611.2365 | 1.2365 | 771.5402 | 161.5402
615 | 318.518 | 0.0039083 | -5.775e-07 | 0 | 318.0225 | -0.49546 |  | 615.0073 | 0.0073438 |  | 614.9662 | -0.033754 | 615.2784 | 0.27844 | 615.7543 | 0.75434 | 782.0314 | 167.0314
620 | 320.1155 | 0.0039083 | -5.775e-07 | 0 | 319.79 | -0.3255 |  | 620.0109 | 0.010853 |  | 619.9643 | -0.03568 | 620.281 | 0.28099 | 620.264 | 0.26404 | 792.6573 | 172.6573
625 | 321.7102 | 0.0039083 | -5.775e-07 | 0 | 321.5575 | -0.15266 |  | 625.0144 | 0.014428 |  | 624.9625 | -0.037455 | 625.283 | 0.28297 | 624.7656 | -0.2344 | 803.4198 | 178.4198
630 | 323.3019 | 0.0039083 | -5.775e-07 | 0 | 323.325 | 0.023075 |  | 630.0181 | 0.018056 |  | 629.9609 | -0.039071 | 630.2844 | 0.28436 | 629.259 | -0.741 | 814.3211 | 184.3211
635 | 324.8908 | 0.0039083 | -5.775e-07 | 0 | 325.0925 | 0.20169 |  | 635.0217 | 0.021723 |  | 634.9595 | -0.04052 | 635.2851 | 0.28514 | 633.7443 | -1.2557 | 825.3629 | 190.3629
640 | 326.4768 | 0.0039083 | -5.775e-07 | 0 | 326.86 | 0.3832 |  | 640.0254 | 0.025415 |  | 639.9582 | -0.041795 | 640.2853 | 0.28529 | 638.2214 | -1.7786 | 836.5474 | 196.5474
645 | 328.0599 | 0.0039083 | -5.775e-07 | 0 | 328.6275 | 0.56759 |  | 645.0291 | 0.029117 |  | 644.9571 | -0.04289 | 645.2848 | 0.28478 | 642.6903 | -2.3097 | 847.8765 | 202.8765
650 | 329.6401 | 0.0039083 | -5.775e-07 | 0 | 330.395 | 0.75487 |  | 650.0328 | 0.032813 |  | 649.9562 | -0.043796 | 650.2836 | 0.2836 | 647.1511 | -2.8489 | 859.3521 | 209.3521
655 | 331.2175 | 0.0039083 | -5.775e-07 | 0 | 332.1625 | 0.94504 |  | 655.0365 | 0.036485 |  | 654.9555 | -0.044508 | 655.2817 | 0.28173 | 651.6038 | -3.3962 | 870.9763 | 215.9763
660 | 332.7919 | 0.0039083 | -5.775e-07 | 0 | 333.93 | 1.1381 |  | 660.0401 | 0.040118 |  | 659.955 | -0.045018 | 660.2791 | 0.27913 | 656.0483 | -3.9517 | 882.7511 | 222.7511
665 | 334.3635 | 0.0039083 | -5.775e-07 | 0 | 335.6975 | 1.334 |  | 665.0437 | 0.043693 |  | 664.9547 | -0.045319 | 665.2758 | 0.27579 | 660.4846 | -4.5154 | 894.6784 | 229.6784
670 | 335.9321 | 0.0039083 | -5.775e-07 | 0 | 337.465 | 1.5329 |  | 670.0472 | 0.047191 |  | 669.9546 | -0.045406 | 670.2717 | 0.27168 | 664.9128 | -5.0872 | 906.7602 | 236.7602
675 | 337.4979 | 0.0039083 | -5.775e-07 | 0 | 339.2325 | 1.7346 |  | 675.0506 | 0.050593 |  | 674.9547 | -0.04527 | 675.2668 | 0.26679 | 669.3328 | -5.6672 | 918.9986 | 243.9986
680 | 339.0608 | 0.0039083 | -5.775e-07 | 0 | 341 | 1.9392 |  | 680.0539 | 0.05388 |  | 679.9551 | -0.044907 | 680.2611 | 0.26108 | 673.7447 | -6.2553 | 931.3955 | 251.3955
685 | 340.6208 | 0.0039083 | -5.775e-07 | 0 | 342.7675 | 2.1467 |  | 685.057 | 0.05703 |  | 684.9557 | -0.044309 | 685.2545 | 0.25454 | 678.1485 | -6.8515 | 943.9529 | 258.9529
690 | 342.1779 | 0.0039083 | -5.775e-07 | 0 | 344.535 | 2.3571 |  | 690.06 | 0.060022 |  | 689.9565 | -0.04347 | 690.2471 | 0.24714 | 682.5441 | -7.4559 | 956.6729 | 266.6729
695 | 343.7322 | 0.0039083 | -5.775e-07 | 0 | 346.3025 | 2.5703 |  | 695.0628 | 0.062834 |  | 694.9576 | -0.042385 | 695.2389 | 0.23886 | 686.9315 | -8.0685 | 969.5574 | 274.5574
700 | 345.2835 | 0.0039083 | -5.775e-07 | 0 | 348.07 | 2.7865 |  | 700.0654 | 0.065444 |  | 699.959 | -0.041048 | 700.2297 | 0.22967 | 691.3108 | -8.6892 | 982.6084 | 282.6084
705 | 346.832 | 0.0039083 | -5.775e-07 | 0 | 349.8375 | 3.0055 |  | 705.0678 | 0.067827 |  | 704.9605 | -0.039453 | 705.2196 | 0.21955 | 695.6819 | -9.3181 | 995.828 | 290.828
710 | 348.3775 | 0.0039083 | -5.775e-07 | 0 | 351.605 | 3.2275 |  | 710.07 | 0.069959 |  | 709.9624 | -0.037595 | 710.2085 | 0.20848 | 700.0449 | -9.9551 | 1009.2182 | 299.2182
715 | 349.9202 | 0.0039083 | -5.775e-07 | 0 | 353.3725 | 3.4523 |  | 715.0718 | 0.071817 |  | 714.9645 | -0.035467 | 715.1964 | 0.19643 | 704.3998 | -10.6002 | 1022.7809 | 307.7809
720 | 351.46 | 0.0039083 | -5.775e-07 | 0 | 355.14 | 3.68 |  | 720.0734 | 0.073374 |  | 719.9669 | -0.033066 | 720.1834 | 0.18337 | 708.7464 | -11.2536 | 1036.5181 | 316.5181
725 | 352.9969 | 0.0039083 | -5.775e-07 | 0 | 356.9075 | 3.9106 |  | 725.0746 | 0.074603 |  | 724.9696 | -0.030385 | 725.1693 | 0.16929 | 713.085 | -11.915 | 1050.4319 | 325.4319
730 | 354.5309 | 0.0039083 | -5.775e-07 | 0 | 358.675 | 4.1441 |  | 730.0755 | 0.075479 |  | 729.9726 | -0.02742 | 730.1542 | 0.15415 | 717.4153 | -12.5847 | 1064.5242 | 334.5242
735 | 356.0621 | 0.0039083 | -5.775e-07 | 0 | 360.4425 | 4.3804 |  | 735.076 | 0.075972 |  | 734.9758 | -0.024167 | 735.1379 | 0.13794 | 721.7376 | -13.2624 | 1078.7971 | 343.7971
740 | 357.5903 | 0.0039083 | -5.775e-07 | 0 | 362.21 | 4.6197 |  | 740.0761 | 0.076056 |  | 739.9794 | -0.02062 | 740.1206 | 0.12063 | 726.0517 | -13.9483 | 1093.2526 | 353.2526
745 | 359.1157 | 0.0039083 | -5.775e-07 | 0 | 363.9775 | 4.8618 |  | 745.0757 | 0.075701 |  | 744.9832 | -0.016776 | 745.1022 | 0.10218 | 730.3576 | -14.6424 | 1107.8926 | 362.8926
750 | 360.6381 | 0.0039083 | -5.775e-07 | 0 | 365.745 | 5.1069 |  | 750.0749 | 0.074878 |  | 749.9874 | -0.01263 | 750.0826 | 0.082588 | 734.6554 | -15.3446 | 1122.7192 | 372.7192
755 | 362.1577 | 0.0039083 | -5.775e-07 | 0 | 367.5125 | 5.3548 |  | 755.0736 | 0.073555 |  | 754.9918 | -0.0081781 | 755.0618 | 0.061813 | 738.945 | -16.055 | 1137.7343 | 382.7343
760 | 363.6744 | 0.0039083 | -5.775e-07 | 0 | 369.28 | 5.6056 |  | 760.0717 | 0.071702 |  | 759.9966 | -0.0034171 | 760.0398 | 0.039833 | 743.2265 | -16.7735 | 1152.9399 | 392.9399
765 | 365.1882 | 0.0039083 | -5.775e-07 | 0 | 371.0475 | 5.8593 |  | 765.0693 | 0.069287 |  | 765.0017 | 0.0016569 | 765.0166 | 0.016623 | 747.4998 | -17.5002 | 1168.338 | 403.338
770 | 366.6991 | 0.0039083 | -5.775e-07 | 0 | 372.815 | 6.1159 |  | 770.0663 | 0.066279 |  | 770.007 | 0.0070473 | 769.9922 | -0.0078449 | 751.765 | -18.235 | 1183.9306 | 413.9306
775 | 368.2072 | 0.0039083 | -5.775e-07 | 0 | 374.5825 | 6.3753 |  | 775.0626 | 0.062643 |  | 775.0128 | 0.012757 | 774.9664 | -0.033595 | 756.022 | -18.978 | 1199.7197 | 424.7197
780 | 369.7123 | 0.0039083 | -5.775e-07 | 0 | 376.35 | 6.6377 |  | 780.0583 | 0.058345 |  | 780.0188 | 0.018789 | 779.9393 | -0.060655 | 760.2709 | -19.7291 | 1215.7073 | 435.7073
785 | 371.2146 | 0.0039083 | -5.775e-07 | 0 | 378.1175 | 6.9029 |  | 785.0534 | 0.053353 |  | 785.0251 | 0.025147 | 784.9109 | -0.089051 | 764.5116 | -20.4884 | 1231.8952 | 446.8952
790 | 372.7139 | 0.0039083 | -5.775e-07 | 0 | 379.885 | 7.1711 |  | 790.0476 | 0.04763 |  | 790.0318 | 0.031832 | 789.8812 | -0.11881 | 768.7441 | -21.2559 | 1248.2856 | 458.2856
795 | 374.2104 | 0.0039083 | -5.775e-07 | 0 | 381.6525 | 7.4421 |  | 795.0411 | 0.041141 |  | 795.0388 | 0.038847 | 794.85 | -0.14996 | 772.9686 | -22.0314 | 1264.8803 | 469.8803
800 | 375.704 | 0.0039083 | -5.775e-07 | 0 | 383.42 | 7.716 |  | 800.0338 | 0.033849 |  | 800.0462 | 0.046193 | 799.8175 | -0.18252 | 777.1848 | -22.8152 | 1281.6813 | 481.6813
805 | 377.1947 | 0.0039083 | -5.775e-07 | 0 | 385.1875 | 7.9928 |  | 805.0257 | 0.025717 |  | 805.0539 | 0.053874 | 804.7835 | -0.21653 | 781.3929 | -23.6071 | 1298.6905 | 493.6905
810 | 378.6825 | 0.0039083 | -5.775e-07 | 0 | 386.955 | 8.2725 |  | 810.0167 | 0.016709 |  | 810.0619 | 0.06189 | 809.748 | -0.25201 | 785.5929 | -24.4071 | 1315.91 | 505.91
815 | 380.1675 | 0.0039083 | -5.775e-07 | 0 | 388.7225 | 8.555 |  | 815.0068 | 0.006784 |  | 815.0702 | 0.070243 | 814.711 | -0.28898 | 789.7847 | -25.2153 | 1333.3416 | 518.3416
820 | 381.6495 | 0.0039083 | -5.775e-07 | 0 | 390.49 | 8.8405 |  | 819.9959 | -0.0040951 |  | 820.0789 | 0.078933 | 819.6725 | -0.32748 | 793.9684 | -26.0316 | 1350.9873 | 530.9873
825 | 383.1287 | 0.0039083 | -5.775e-07 | 0 | 392.2575 | 9.1288 |  | 824.984 | -0.015969 |  | 825.088 | 0.087962 | 824.6325 | -0.36754 | 798.1439 | -26.8561 | 1368.849 | 543.849
830 | 384.6049 | 0.0039083 | -5.775e-07 | 0 | 394.025 | 9.4201 |  | 829.9711 | -0.028877 |  | 830.0973 | 0.097331 | 829.5908 | -0.40917 | 802.3112 | -27.6888 | 1386.9287 | 556.9287
835 | 386.0783 | 0.0039083 | -5.775e-07 | 0 | 395.7925 | 9.7142 |  | 834.9571 | -0.04286 |  | 835.107 | 0.10704 | 834.5476 | -0.45242 | 806.4705 | -28.5295 | 1405.2282 | 570.2282
840 | 387.5488 | 0.0039083 | -5.775e-07 | 0 | 397.56 | 10.0112 |  | 839.942 | -0.05796 |  | 840.1171 | 0.11709 | 839.5027 | -0.4973 | 810.6215 | -29.3785 | 1423.7494 | 583.7494
845 | 389.0164 | 0.0039083 | -5.775e-07 | 0 | 399.3275 | 10.3111 |  | 844.9258 | -0.074219 |  | 845.1275 | 0.12747 | 844.4562 | -0.54385 | 814.7644 | -30.2356 | 1442.4943 | 597.4943
850 | 390.4811 | 0.0039083 | -5.775e-07 | 0 | 401.095 | 10.6139 |  | 849.9083 | -0.09168 |  | 850.1382 | 0.1382 | 849.4079 | -0.59209 | 818.8992 | -31.1008 | 1461.4647 | 611.4647
[7 empty] | [-200, 850] | average | 0.013699 |  | average | 0.010351 | average | 0.0085458 | average | 0.0012476 | average | 99.539
[8 empty] | max | 0.13086 |  | max | 0.23158 | max | 0.49895 | max | 15.2554 | max | 611.4647
[8 empty] | min | -0.23939 |  | min | -0.51347 | min | -1.3536 | min | -31.1097 | min | -0.0035417
[7 empty] | [-150, 750] | average | 0.016968 |  | average | 0.014784 | average | 0.069702 | average | 4.0448 | average | 61.7678
[8 empty] | max | 0.13086 |  | max | 0.23158 | max | 0.49895 | max | 15.2554 | max | 372.7192
[8 empty] | min | -0.088036 |  | min | -0.12493 | min | -0.25745 | min | -21.2554 | min | -0.0035417
222 | 183.9181 | 0.0039083 | -5.775e-07 | 0 | 100 |  |  | 222.0727 | 0.072667
223 | 184.2832 | 0.0039083 | -5.775e-07 | 0 | 100 |  |  | 223.0729 | 0.07288
224 | 184.6483 | 0.0039083 | -5.775e-07 | 0 | 100 |  |  | 224.0731 | 0.073084
226.66 | 185.6186 | 0.0039083 | -5.775e-07 | 0 | 100 |  |  | 226.7336 | 0.073575
86.235 | 133.2738 | 0.0039083 | -5.775e-07 | 0 | 100 |  |  | 86.1968 | -0.038179